## Slide 1: Run2-TDC Dataformat
Su-Yin Grass Wang, Jin-Yuan Wu, David Christian Da-Shung Su and Wen-Chen Chang

## Slide 2: Register Map
0x000 version number (e9061212)
0x004 CSR
0x008 test area =0xaaaaaaaa
0x00c
  BLKS(26-24) |DeMultiDD(23-16) |CR_Reset(15) | CR_Clear(0)
0x010 0xe906e0## = # of events
0x014
  Time Window On(31)
  UP_LIMIT(25-16) | LOW_LIMIT(9-0)
0x030 –Time window Check
0x1000~0x1ffc = event buffer (1k internal)
0x2000~0x2ffc = test area
D31-D16 | D08 | D7-5 | D2-0
EN4X (CSR) | BothEdge (Edge selection) | SEL(READ/WRITE addr) | M (buffer size)

## Slide 3: Configuration
Event table
  Adjustable
    Time window width
    CIP time
CSR [7:5] | Number of Circular Buffers | Circular Buffer Size (MB) | Buffer Timing Window | CIP Time
000 | 8 | 32 | 512 ns (-256 ns) | 32 x 16 x 16 ns = 8 us
001 | 4 | 64 | 1024 ns (-256 ns) | 64 x 16 x 16 ns = 16 us
011 | 2 | 128 | 2048 ns (-256 ns) | 128 x 16 x 16 ns = 32 us
CSR [2:0] | Number of events | Max. number of words/event
5 | 32 | 32
4 | 16 | 64
3 | 8 | 128
2 | 4 | 256
1 | 2 | 512
0 | 1 | 1024
Event Buffer
  Adjustable
    buffer size
    number of buffer
Notes: These buffer size corresponding to hit accepting window of 512, 1024 and 2048 ns. Due to effect of the RegPipe4a, the 100% hit accepting windows are 256 ns narrower.
The 256 ns here is the part close to the trigger of TDC. Since our trigger is delayed by 700us, we may not care.

## Slide 4: Scalar
TC[2]
62
63
0
1
2
3
4
5
TC[8..3]
62
63
0
1
2
3
4
5
TCQ[2]
TCQ[8..3]
SE[0]
SE[1]
CH0
1
2
3
4
5
QS[7..0]
SEL=TC[2]
62
63
0
1
2
3
4
5
ADDRA[5..0]
=TC[8..3]
(for read)
62
63
0
1
2
3
4
5
ADDRA[5..0]
(for write)
Write
(inv. RWA)
=TC[2]
8 buffers for scalar
0x00c , BLKS(26-24)= buffer (0~7)

## Slide 5: Multiple Hit Elimination
TFQ[0] (Data valid)
Reg1_0_Q[0]
DFN1_1_Q
(DFN1_2Off if DD[6]=1)
DFN1_0_Q
(DFN1_2On if DD[6]=1)
DFN1_2_Q
3
3
2
1
0
3F
3E
3D
X
Q[5:0]
3C
3
DFN1_3_Q
Hits from Reg1_0_Q[0] are blocked in this region.
DD is the register controls multiple hits elimination.
DD[7]
  =1: Eliminating Multiple Hit
DD[6]
  =1: Updating Mode
  =0: Non-updating Mode
DD[5:0]
  0-63: Number of cycles of multiple hit elimination. 4ns*(4+DD[5:0])
  (updating mode 0-31 only)

## Slide 6: Multiple Hit Elimination
In register map : 0x00c
  BLKS(26-24) |DeMultiDD(23-16) |CR_Reset(15) | CR_Clear(0)
  DeMultiDD (5-0) length of Inhibit : 0-63: Number of cycles of multiple hit elimination. 4ns*(4+DD[5:0]) (16 – 272 ns)
  (updating mode 0-31 only)
  DeMultiDD (6) update mode/NON –update mode
  DeMultiDD (7) Elimination on/off
Raw data
Inhibit
TDC recorded

## Slide 7: DATA format
D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D0
Trig [13:4] (trigger time) |  |  |  |  |  |  |  |  |  |  |  | Fine time |  |  | 
D31 | D30 | D29 | D28 | D27 | D26 | D25 | D24 | D23 | D22 | D21 | D20 | D19 | D18 | D17 | D16
1 | CNT[10:0] (number of hits) (could indicate truncated data readout ) |  |  |  |  |  |  |  |  |  |  | 0 | 0 | 0 | 1
D31 | D30 | D29 | D28 | D27 | D26 | D25 | D24 | D23 | D22 | D21 | D20 | D19 | D18 | D17 | D16
Cable: 4,5,6,7 |  |  |  | Channel number: 0-15 |  |  |  | hit number |  |  |  | 0 | 0 | 0 | positive or negative edge
Header
Data
D15 | D14 | D13 | D12 | D11 | D10 | D09 | D08 | D07 | D06 | D05 | D04 | D03 | D02 | D01 | D0
TDC [13:4] (data time) |  |  |  |  |  |  |  |  |  |  |  | Fine time |  |  | 
Fine time = 0~8, ~444ps
TDC = 4 ns
Notes: MSB = most significant bit

## Slide 8: CODA raw data – FEE information event
0x84 (event type)
  (0x84 may already used, could change to whatever more than 128 but not used)
0xe906f011 = FEE ID
0xffff0060 =TDC Reg(0x004)
0x00840000 =TDC Reg(0x00c) (Multihit Elimilation)
0x82000001 =TDC Reg (0x014) (Time window)

## Slide 9: CODA raw data – physics event
Ignore every 0xe906e906 in DATA
header
Data
boardID/ Num of word in TDC buffer
FEE type
Dummy word